Appendix F: Summary of Station-Level Findings and Conclusions

| Zone | Station | Temperature |  | Relative Humidity | Atmospheric Pressure | Rainfall | Overall station conclusion |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Western – Trans Nzoia High Rainfall | Eldoret | Good agreement. Tmax strongest, r = 0.723; all three temperature statistics met Threshold. |  | Strong temporal agreement, particularly RH Mean (r = 0.912), but Below Threshold because of larger absolute differences. | Weak agreement, r ≈ 0.32–0.33; Below Threshold. | Useful temporal agreement; RG1 r = 0.790, the highest rainfall correlation in the network. | Good temperature and rainfall temporal performance. Humidity reproduces variability well but requires calibration; pressure requires further investigation. |
| Eastern – Marsabit/Garissa/Tana River Pastoral | Garissa | Strongest temperature performance. Tmax r = 0.940, RMSE = 0.776°C; Breakthrough. Tmin also Breakthrough; Tmean Threshold. |  | Strong temporal agreement, but absolute differences resulted in Below Threshold classifications. | Variable agreement; Mean Pressure r = 0.671, but pressure classifications were Below Threshold. | High coverage (99.1%) with useful temporal comparison; both gauges assessed. | Strongest overall validation case, particularly for temperature, supported by the highest mean station coverage (90.7%). |
| Western – Trans Nzoia High Rainfall | Kitale–New | Limited evidence. Correlations ranged from r = 0.379–0.616; all Below Threshold. |  | Temporal relationships present, but low coverage limits confidence. | Weak agreement and large errors; e.g. Pressure Max r = 0.132, RMSE = 18.280 hPa. | Very limited paired coverage (21.7–25.2%); RG1 r = 0.019. | Interpret cautiously. Very low paired-observation coverage substantially limits the strength of the validation evidence. |
| South East – Makueni/Kitui Marginal-Agricultural | KITUI |  | Moderate agreement; Tmean r = 0.533; Tmin showed RMSE = 2.810°C. | Not included in formal assessment – insufficient eligible data/assessment coverage for inclusion in the consolidated results. | Not included in formal assessment – insufficient eligible data/assessment coverage for inclusion in the consolidated results. | Good rainfall coverage for RG1/RG2; RG1 r = 0.636 and coverage up to 92.8%. | Rainfall provides useful evidence, while limited parameter coverage restricts conclusions for humidity and pressure. |
| Coastal – Kilifi/Taita-Taveta Coastal | MSABAHA AGROMET |  | Strong performance. Tmean r = 0.829, RMSE = 0.889°C; Breakthrough. | Good temporal agreement, but Below Threshold because of absolute differences. | Weak temporal agreement; Pressure Max r = 0.285; Below Threshold. | Good coverage; RG1 coverage 94.7%, with useful internal gauge comparison. | Strong temperature performance, with rainfall providing useful additional evidence. Humidity and pressure require refinement. |
| Coastal – Kilifi/Taita-Taveta Coastal | MTWAPA AGROMET |  | Strong temporal agreement. Tmean r = 0.869 and RMSE = 0.760°C; Breakthrough. | Moderate temporal agreement; RH Mean r = 0.714, but Below Threshold. | Highly data-limited. Pressure coverage falls to 0.7% for some assessments and results require caution. | Good rainfall coverage; RG1 86.1%, RG2 86.1%; both Below Threshold under the implemented criterion. | Strong temperature result, but the low pressure coverage demonstrates the importance of considering data availability alongside statistical performance. |
| South Rift – Narok Agro-Pastoral | NAROK |  | Good agreement; Tmin r = 0.892 and Tmax r = 0.836; all temperature results Threshold. | Strong temporal agreement; RH Mean r = 0.891, although Below Threshold on absolute accuracy. | Best pressure temporal relationships among several stations; Pressure Max r = 0.794, but Below Threshold. | Good rainfall performance and coverage; RG1 r = 0.618, RMSE = 5.325 mm, coverage 94.6%. | Balanced performance across temperature, humidity and rainfall, with pressure remaining the principal area requiring improvement. |
| Central – Nyandarua/Nyahururu High Altitude | NYAHURURU AGROMET |  | Good temporal agreement; Tmin r = 0.771; temperature results generally Threshold. | Strongest humidity correlation in the network for RH Mean, r = 0.923, but Below Threshold because of absolute differences. | Not included in formal assessment – the available consolidated result set did not contain sufficient eligible pressure comparison results. | Good rainfall agreement; RG1 r = 0.697, coverage 82.5%. | Strong temporal reproduction of humidity, temperature and rainfall, although humidity absolute bias and pressure data availability require attention. |