Appendix l: 3DPAWS Maintenance Manual

# 3D-PAWS MAINTENANCE MANUAL

# INTRODUCTION

3D-PAWS is a cost-effective automatic weather station designed to measure:

- Rainfall

- Temperature

- Humidity

- Pressure

- Wind speed & direction

- UV

It uses 3D-printed structures integrated with sensors and a microcontroller for real-time data transmission.

Due to environmental exposure, regular maintenance is mandatory to ensure accuracy and reliability.

This manual provides detailed maintenance procedures for 3D-Printed Automatic Weather Stations (3D-PAWS). It is tailored to the actual hardware configuration including dual rain gauges, 3D-printed wind sensors, and modular PVC mounting systems.

The targeted user of this manual is the 3D PAWS site caretaker who should liaise with a 3D PAW technically competent person during the time of the system maintenance.

# PREPARATION FOR THE MAINTENANCE

The field maintenance requires a few very basic tools as listed here below;

- A pliers

- Spirit level

- flat screw driver (A bout 5mm head)

- A star head screw driver (PH2 or PH3 size)

- Soft brush (Tooth brush will work better)

- Water (In a bucket)

Prepare a check list for all the work intended to be carried out like the example given here below;

## Checklist sample

| Component | Check | Status |
| --- | --- | --- |
| Rain Gauge | Clean & tipping |  |
| Radiation Shield | Clean & airflow |  |
| Wind Sensors | Free rotation |  |
| Electronics | Dry & active |  |
| Solar Panel | Clean & charging |  |
| Structure | Stable & upright |  |

## Contact the 3 D PAWS Technician

When planning to carry out the routine maintenance, it is necessary to contact a 3 D PAWS technician a day before. The technician may offer necessary assistance especially in case of network issues.

# SYSTEM COMPONENTS

- Dual Rain Gauge System (Funnel, Mesh, Tipping Bucket, Magnet)

There are two such rain gauges per station

- Radiation Shield (Multi-plate, Temperature/Humidity Sensor)

- Wind Sensors (3D-printed Anemometer and Vane)

Wind direction sensor

Wind speed sensor

- Electronics (Boron Logger, Battery, Wiring)

- Solar Power System

- Structural Frame (PVC + metal mast)

# MAINTENANCE SCHEDULE

| Activity | Frequency |
| --- | --- |
| Clean solar panel | Weekly |
| Clean rain gauge funnel & mesh | Weekly |
| Inspect tipping bucket | Monthly |
| Clean radiation shield | Monthly |
| Inspect electronics (moisture/insects) | Monthly |
| Check alignment & structure | Monthly |
| Grass cutting / site clearing | As needed |

# RAIN GAUGE MAINTENANCE (DETAILED)

Step 1: Remove mesh filter and rinse with clean water. Use a soft brush to remove debris.

Mesh

Filter

Step 2: Remove funnel gently by loosening clamp and bolt and lifting vertically

Bolt

Clamp

Step 3: Inspect funnel neck for blockage. Clear using water or thin rod.

Make sure this hole is clean

Step 4: Open tipping bucket housing by removing clips on both sides.

Clips

Step 5: Carefully remove tipping bucket without bending pivot.

Pivot

Step 6: Clean bucket chambers and pivot axis. Ensure zero friction.

Buckets

Step 7: Inspect magnet at the back of the bucket (Report if missing or highly corroded)

Magnet

Step 8: Reinstall bucket ensuring correct alignment with sensor.

Sensor

Step 9: Reassemble housing and funnel while making sure that the base is level by use of a spirit level provided.

Spirit level

Step 10: Perform test by pouring water slowly and confirming consistent tipping.

Pour water here to test the tipping

# RADIATION SHIELD MAINTENANCE

Step 1: Inspect for dust accumulation between plates.

Step 2: Use cloth or compressed air to clean.

Step 3: Check for insect nests and remove.

Step 4: Ensure plates are evenly spaced and not warped.

# WIND SENSOR MAINTENANCE

Step 1: Manually spin cups; verify free rotation.

Step 2: Inspect for cracks or imbalance.

Step 3: Check vane movement and alignment.

Step 4: Inspect shaft for friction; clean if necessary.

Step 5: Tighten mounting connections.

# ELECTRONICS MAINTENANCE

Step 1: Open data logger enclosure box carefully to avoid falling parts which are not fitted on the enclosure case.

Data logger battery

Data logger

Main battery connected to the solar panel

Gasket (Prevent water and dust from getting in)

Step 2: Remove insects and debris.

Heavy buildup of insects

Step 3: Check battery condition and wiring.

Insects affected data logger battery

Step 4: Ensure enclosure is dry. Check the gasket (rubber around the closing point of the data logger box. Either the gasket is poorly closed or worn out

Water that has found its way into the data logger box through the gasket

Gasket

Step 5: Check if there is any light on the Boron data logger. If there is no light, this may mean that the battery is bad

If the Boron data logger is healthy, there should be a blue flickering light here

Step 6: Report any faulty components.

# FIELD CHECKLIST

After everything is done, fill the check list appropriately as indicated below. File the check list for future reference.

| Component | Check | Status | Comments |
| --- | --- | --- | --- |
| Rain Gauge | Clean & tipping | Ok | The base was out of level but leveled |
| Radiation Shield | Clean & airflow | Ok |  |
| Wind Sensors | Free rotation | OK |  |
| Electronics | Dry & active | Ok |  |
| Solar Panel | Clean & charging | Ok |  |
| Structure | Stable & upright | Ok |  |

# TROUBLESHOOTING LOGIC

## RAIN FAILURE:

Check funnel → Clean → Test

If still faulty → Check bucket movement

If stuck → Clean pivot

If not → Either the magnet or faulty wiring (Report the issue)

## WIND FAILURE:

Check rotation → If stuck → Clean

If free → Check cable

If OK → Report the issue

## SYSTEM OFFLINE:

Check battery → Report if dead

Check solar → Clean panel

Check connectivity → Report

# DESIGN-SPECIFIC NOTES

Due to 3D-printed components, periodic inspection for wear, cracks, and UV degradation is essential. Dual rain gauge systems require calibration comparison to detect faults.